외자구매 입찰공고

입찰공고 제06-05호

1. 입찰에 부치는 사항

번호	품 명	수량	입찰참가 신청마감	입찰일시	입찰방법
1	Nano Indentation System (나노 압입 시험기)	1 set	2006.5.02 14:00	2006.05.02 14:30	규격, 가격 분리입찰

2. 입찰참가자격

-당 연구원 외자구매요령 제11조에 의한 유자격자

3. 낙찰자 선정방법

-규격서 적격자가중 Target price내의 최저가  단, 규격적격자가 2인이상인 경우로서 개찰결      과 Target price이하로 입찰한 자가 없을 때에는 가격입찰을 할 수 있음.

-규격서는 한국기계연구원 요청규격에 의함

4. 입찰장소 : 당 연구원 입찰실

5. 제출서류

①입찰참가신청서 ②사업자등록증 ③입찰보증금 또는 이행보증보험증권(입찰금액의 100분의

5이상) ④인감증명원 및 사용인감계 ⑤위임장 ⑥법인등기부등본 ⑦국세, 시세 완납증명서

⑧Catalogue 2부 ⑨무역대리점 계약서 ⑩규격입찰서 2부, 가격입찰서 1부

⑪Maker's Certificate(원본)

-상기 구비서류를 입찰일 참가신청 마감시간까지 반드시 제출.

-가격입찰(F.O.B기준)대상자는 규격적격자로 확정된 업체에 한하여 별도 통지함.

-가격 평가기준은 입찰일의 최초로 고시되는 우리은행 전신환 매도율임.

(원단위 미만 금액은 절삭)

6. 입찰보증금 및 본원귀속

-낙찰자가 소정 기일 내에 계약을 체결하지 아니한 경우, 입찰보증금은 국가를 당사자로

하는 계약에 관한 법률시행령 제38조에 의해 본원에 귀속되며 부정당업자로 제재를 받게

됩니다.

7. 입찰무효

-국가를 당사자로 하는 계약에 관한 법률시행령 제39조 제3항, 동 시행규칙 제44조에 해당

되는 입찰은 무효입니다.

8. 기타사항

-기타 입찰에 관한 사항은 우리 연구원(창원) 시설자재팀(055-280-3724)으로, 기술적인 사항      은 재료안전연구센터(055-280-3677)로 문의하여 주시기 바라며 본 공고문 및 규격서 내용은       우리 연구원 홈페이지(http://emille.kimm.re.kr)를 통해서도 볼 수 있습니다.

위와 같이 공고합니다.

2006. 04. 20.

한 국 기 계 연 구 원 장

입 찰 참 가 신 청 서 ※아래사항 중 해당되는 경우에만 기재하시기 바랍니다.
신청인	상호또는법인명		사업자등록번호
	주 소		전 화 번 호
	대 표 자		주민등록번호
입찰개요	입 찰 공 고 (지명) 번 호		입 찰 일 자
	입 찰 건 명
입 찰 보 증금	납 부	․보증금율 : % ․보증금액 : 금 원정(￦ ) ․보증금납부방법 :
	납 부 면 제 및 지 급 확 약	․사유 : ․본인은 낙찰후 계약미결시 귀원에 낙찰금액에 해당하는 소정의 입찰보증금을 현금으로 납부할 것을 확약합니다.
대리인사용인감	본 입찰에 관한 일체의 권한의 다음의 자에게 위임합니다. 성명 : 주민등록번호 :		본 입찰에 사용할 인감을 다음과 같이 신고합니다. 사용인감 (인)
본인은 위의 번호로 공고(지명통지)한 일반(제한․지명)경쟁입찰에 참가하고자 귀원에서 정한 공사(물품구매․기술용역)입찰유의서 및 입찰공고사항을 모두 승낙하고 별첨 서류를 첨부하여 입찰참가 신청을 합니다. 붙임서류 : 1. 입찰서 년 월 일 입찰참가신청자 (인) 한국기계연구원장 귀하

영      문     규     격     서

COMMODITY DESCRIPTION

H.S.NO.	DESCRIPTION	UNIT	QUANTITY
9024.10.9000	Nano Indentation System	System	1

A. Features:

-	This instrument should be performed Young’s modulus, Hardness, Fracture toughness, Stress exponent/strain-rate sensitivity, Stroage and loss modulus, True coefficient of friction, Scratch hardness, Scratch critical load and Quantitative surface profiles at the nano scale.

-	This instrument should be perfect for measuring virtually all material types, including Metal, Thin films, Multi-phase material, Compositionally modulated materials, Ceramics, Intermetallic alloys, Ion-implanted materials, Heat treated materials, Polymer Composites, Surface treated materials, Biological materials.

-	The instrument shall allow quantitative measures of hardness and modulus to be made during the loading cycle.

-	By using the continuous stiffness measurement technique, the stiffness data along with load and displacement data dynamically, allowing hardness and Young’s modulus to be calculated at every data point acquired during the indentation experiment.

-	Fully automated, computerized data acquisition and control system including monitor and keyboard.

-	It should include Vibration isolation and environmental enclosure.

-	Certified by ISO 9001.

-	The vendor shall provide all technical data and catalog for reviewing the bid proposal.

B.	Specifications

1. Indentation Head Assembly

-	Displacement resolution : < 0.01 nm

-	Total indenter travel : 2 mm or more

-	Maximum indentation depth : > 500 μm

-	Load application : Coil/magnet assembly

-	Displacement measurement : Capacitance gauge

-	Maximum load : 500 mN or more

-	Maximum load should be available up to 1 kgfwith high –load option

-	Load Resolution : better than 50 nN

-	Conatct force : below than 1.0μN

-	Berkovich diamond indentation tip.

2. Indentation Placement

-	Useable surface area : 90 mm x 100 mm or more

-	Motorized x-y sample manipulation table with computer mouse control

-	Position accuracy : 1.5 μm or better

3. Continuous Stiffness Measurement

-	Continuous force oscillation, frequency 1 to 300 Hz

-	Force amplitudes of 0.1μN to 4.5 mN

-	Software with dynamic model characterizing the instrument’s dynamic response in the frequency range of 0.1 to 2000 Hz

-	System controls, monitors and analyzes force and displacement oscillation amplitude oscillation.

-	Feedback control of the force oscillation to maintain a user-defined displacement amplitude oscillation.

-	Able to measure contact stiffness with displacement oscillation amplitudes as small as 1 angstrom.

4. Microscope

-	Color optical imaging system include video microscope image on the computer screen with the system software.

-	Microscope : 10X, 40X, 1000X using CCD.

5. Vibration isolation system

-	Self-tuning negative stiffness vibration isolation table

-	Passive environmental enclosure

-	Resonant frequency should be equal or less than 0.5 Hz

6. Nano Positioning

-	Incorporates a nano positioning stage with controller that is fully-integrated with the Nano Indenter system software.

-    Unique design enables simple changeover from either the Nano Indenter system's standard performance micro positioning tables or high performance micro positioning tables to the nano positioning tables. Includes the following specifications for the positioning stage:

-    100 μm x 100 μm range in X and Y

-    Positioning accuracy of 2nm

-    Linearity error less than 0.02%

-    120 Hz to 350 Hz resonant frequency (depending on the load)

-    Features full integration with TestWorks software.

-   Specific task built,allowing parallel control of the positioning stage with the nano indentation head (load, displacement, and frequency-specific measurement)

-     Data recorded just as other channels in TestWorks®

7. Heated Sample Stage up to 350C

-  This option includes provision of a local heating stage to raise the sample to a maximum temperature of 350ºC.The custom sample heating stage option encompasses the following items:

-	Custom temperature controller with two selectable heating ranges R.T to 200C and 200C to 400C

-	Argon flow around the sample to avoid corrosion due to heat

-	Heat shield to protect the indenter head and to limit the heat in the indentation head

-	Liquid cooling system for waste heat removal

Lead time for this item is approximately 120 days

MTS cannot guarantee the validity of results when using a heated sample stage due to possible temperature differences between the sample stage and the sample surface.

8. Dynamic Contact Module

-	DCM as an additional indentation head to the Nano Indenter (R) XP system. Specifications for the Dynamic Contact Module indentation head are as follows:

-	Load resolution 1nN.

-	Displacement resolution 0.0002nm.

-	Maximum load 10mN.

-	Maximum indentation depth/maximum z travel >15um.

-	AccuTip (TM) Berkovich diamond indentation tip.

9. Berkovich diamond Tip

-	an additional set of Berkovich tip

10. Cube Corner diamond Tip

-	one set of Cube Corner diamond Tip

11. Computer

-	Pentium 4 2.4 GHz, 256 Mb RAM, 40 Gb hard drive, 48x Max CDRW drive, 3.5 inch 1.44 Mb floppy drive.  17 inch LCD flat panel display, integrated network interface port 10/100/1000, integrated Intel video DVMT, integrated SoundBlaster compatible sound, color inkjet printer.

12. Software

-	Windows operating system

-	It shall be enable to share data easily across your network and export it to virtually any application for analysis and display.

-	Control of loading types and rates

-	Control of load-hold length

-	Control of data logging rates

-	Control of unloading types and rates

-	Control of sequence and number experimental segments

-	PID control of any calculated channel.

-	For test setup and examination, it shall have a view of the overall sample mount for convenient motion from one sample to any other, and another view of the sample surface through the video microscope.

-	Forreview, it shall be able to examine the results of the tests for any or all indentations made during the test.

-	For experiment set-up, the define view shall graphically display each step of the test process and include a " wizard" for stepping through the experiment setup process.

-	It shall be set up a test batch with as many as 500 indents.

-	Multiple tests on multiple specimen should be possible.

-	All of the indent locations shall be selected individually or in arrays using the on-screen video microscope.

-	Thetest data shall be automatically stored as Microsoft Excel files, making the automated report generator convenient to use and modify to fit  user style.

-	Standard calculation shall included shall be calculation of hardness and modulus using the Oliver & Pharr method, calculationof hardness and modulus as a function of depth and /or storage and loss modulus as well as other more specialized calucaltion.

-	It shall be able to set up the software to use own method for calculation tip geometry and area function.

-	It shall be able to set up the software to use own method for analytical solutions to the problem of pile-up.

C. Remarks

-	Installation support and check-out service shall be provided by supplier's engineer.

-	On-line system manuals.

-	All expenses for Factory training of 2 persons should be provided by supplier.

-	One year system warranty after acceptance.

-	System operational training should be provided by supplier.

-	The bidder must note that requirements in the above specifications are intended to be descriptive only and not restrictive. The bids offered under other specifications will be considered, provided that the bid clearly states the offered commodities are substantially superior to those requirements in the above specification. (by descriptive literature, illustrations, etc)